# DescSpim Deepsky Construction manual

Steven Edwards, PhD, Science for Life Laboratory

Maximilian Senftleben, PhD, Uppsala University

# Table of Contents

Table of Contents	1

1.	Microscope Construction	2

Assembling the detection arm stage components	2

Assembling the detection objective and filter slider	2

Assembling the tube lens and camera	3

Mounting the detection arm to the stage	3

Assembling the sample stage	4

Assembling the illumination arm stage components:	4

Assembling the illumination cage components	5

Laser, power supplies and triggering	5

# Microscope Construction

## Assembling the detection arm stage components

- Obtain a crossed-roller side-actuated translation stages (XR25P/M, Thorlabs), loosen the flexure lock, and remove the micrometer head.

- If necessary, convert the stage to a left-hand configuration by swapping the side mounts with the locking plate.

- CRITICAL STEP Fully retract the micrometer head before removing the locking plate to prevent the stage from rapidly retracting.

- Mount one motorized actuator (Z825B, Thorlabs) onto the translation stage, labeling the assembly as the camera stage

- Fix the stage to the to the optical breadboard at the given location (Fig x) with M6 cap screws

- Attach a 66 mm dovetail rail (XT66SD-250, Thorlabs) to the translation stage with M6 cap screws.

- Mount 3 rail carriages (XT66P2/M, Thorlabs) on the 66 mm dovetail rail.

- Fix three post holders (PH30/M, Thorlabs) to the rail carriages with M6 setscrews

## Assembling the detection objective and filter slider

- Attach the W26 x 0.706 internal to SM1 external thread adapter (SM1A27, Thorlabs) to the detection objective (MY10X-803, Thorlabs).

- Attach the objective thread adaptor to the SM1 internal to SM2 external thread adapter (SM2A6, Thorlabs).

- Attach the SM2 external thread adaptor to the 60 mm cage plate, 0.5” thick (LCP34/M Thorlabs).

- Insert two 2” long cage assembly rods (ER2, Thorlabs) in the bottom two holes of the 60 mm cage plate and tighten the M4 set screws.

- Attach a 40 mm optical post (TR40/M, Thorlabs) to the cage plate using an M4 set screw.

- Attach the SM1-threaded 30 mm to 60 mm cage plate adapter, 0.5” thick (LCP33/M, Thorlabs) via the bottom two holes to the cage assembly rods and tighten the M4 set screws.

- Insert four 1” long cage assembly rods (ER1, Thorlabs) into the centre four holes of the SM1-threaded 30 mm to 60 mm cage plate adapter.

- Insert the filter slider (ELL9, Thorlabs) between the 60 mm cage plate and SM1-threaded 30 mm to 60 mm cage plate adapter, fastening to the four previously inserted 1” long cage assembly rods.

- Attach a 40 mm optical post (TR40/M, Thorlabs) to the SM1-threaded 30 mm to 60 mm cage plate adapter with an M4 set screw.

## Assembling the tube lens and camera

- In order, screw together the following components: 30 mm cage plate (CP33/M, Thorlabs), SM1 external to SM2 internal thread adapter (SM1A2, Thorlabs), 0.5” SM2 lens tube (SM2L05, Thorlabs), f = 200 mm tube lens (TTL200-A, Thorlabs), 3” SM2 lens tube (SM2L30, Thorlabs), 0.5” SM2 lens tube (SM2L05, Thorlabs), SM1 external to SM2 internal thread adapter (SM1A2, Thorlabs), 30 mm cage plate (CP33/M, Thorlabs).

- Insert four 1” cage assembly rods into the second 30 mm cage plate and tighten the retaining screws.

- Fasten a second 30 mm cage plate onto the cage assembly rods, leaving a small gap between the two.

- Screw a SM1 coupler (SM1T10, Thorlabs) onto the second 30 mm cage plate.

- Screw a 0.5” SM1 lens tube (SM1L03, Thorlabs) onto the SM1 coupler.

- Screw the Kiralux 8.9 MP Monochrome CMOS camera (CS895MU, Thorlabs) onto the 0.5” SM1 lens tube.

- Retaining ring.

CRITICAL STEP The assembly allows for the camera to be rotated to the vertical position to maximise the field-of-view.

CRITICAL STEP The precise alignment of the camera and tube lens is critical for diffraction limited imaging. The distance between the camera and tube lens can be adjusted by loosening the retaining screws on the 30 mm cage plate closest to the camera and moving the camera attached to the cage plate along the optical path direction. To ensure optimal alignment, point the camera outside the window and align the camera by focusing on structures 50 meters away.

## Mounting the detection arm to the stage

- Attach the detection objective and filter slider to the camera and tube lens assembly using the four 1” long cage assembly rods and tighten the retaining screws.

- The complete detection arm can now be attached to the detection arm stage using the three 40 mm optical posts and post holders.

CRITICAL STEP Ensure the detection arm is level before fixing in place using a circular bullseye level (LVL03, Thorlabs).

## Assembling the sample stage

- Fasten a 3” dovetail rail (XR25DR2, Thorlabs) flush with the corner of the optical breadboard using M6 set screws.

- Obtain a crossed-roller side-actuated translation stages (XR25P/M, Thorlabs), loosen the flexure lock, and remove the micrometer head.

- If necessary, convert the stage to a left-hand configuration by swapping the side mounts with the locking plate.

CRITICAL STEP Fully retract the micrometer head before removing the locking plate to prevent the stage from rapidly retracting.

- Mount one motorized actuator (Z825B, Thorlabs) onto the translation stage, labeling the assembly as the sample stage.

- Mount the translation stage onto the dovetail rail.

- Fix two 25 mm pillar posts (RS25/M, Thorlabs) to the back two counterbored M6 holes on the underside of the Self-Contained XYZ 25 mm Translation stage (LX30/M, Thorlabs) using M6 cap screws.

- Fix two 25 mm pillar posts (RS25/M, Thorlabs) to the front two M6 counterbored holes of the translation stage using M6 set screws.

- Fasten the Self-Contained XYZ 25 mm Translation stage to the two posts on the front of the translation stage using M6 cap screws.

- Create a ledge for the cuvette using a mounting base BA1/M attached to the vertical translational stage using an M6 cap screw.

- Fasten a microscope slide spring clip to the vertical translation stage using an internal M4 x 0.7 threads and external M6 x 1.0 threads adapter (AE4M6MS, Thorlabs).

## Assembling the illumination arm stage components:

- Obtain two crossed-roller side-actuated translation stages (XR25P/M, Thorlabs). From one translation stage, loosen the flexure lock, and remove the micrometer head.

- If necessary, convert the stage to a right-hand configuration by swapping the side mounts with the locking plate.

CRITICAL STEP Fully retract the micrometer head before removing the locking plate to prevent the stage from rapidly retracting.

- Fix the stages to the to the optical breadboard at the given location (Fig x) with M6 cap screws.

- With the micrometer fully retracted fasten a 500 mm long, 66 mm wide single dovetail rail (XT66SD-500, Thorlabs) to the translation stages with four M6 cap screws.

- Mount 4 rail carriages (XT66P2/M, Thorlabs) on the 66 mm dovetail rail.

- Fix three post holders (PH30/M, Thorlabs) to the front three rail carriages with M6 setscrews.

- Fix one ⌀12.7 mm pedestal post holder (PH30E/M, Thorlabs) to the rear rail carriage with a clamping fork (CF038-P5, Thorlabs) and M6 cap screw.

## Assembling the illumination cage components

- Insert four cage rods (ER4, Thorlabs) into a right-angle kinematic elliptical mirror mount (KCB1EC/M, Thorlabs) and tighten the retaining screws.

- Slide a 30 mm cage system adjustable slit (VA100CP/M, Thorlabs) onto the cage rods with the slit orientated XX.

- Insert collimating lens XX into a XX lens mount and secure with a retaining ring.

- Insert fiber XX into a 30 mm cage plate (CP33/M) and slide onto the cage rods, fastening the retaining screws

- Insert mirror.

- Temporarily insert optical fibre and collimate the beam path using a shear plate.

TO CONTINUE

## Laser, power supplies and triggering

- Route each stage’s motor cable to its KDC101 controller (Thorlabs), plug each controller into its KSP101 power supply (Thorlabs), and arrange the controller–power pairs neatly on the breadboard.

- The stages, filter slider lasers are triggered using an Arduino Uno

# Sample preparation

# Hardware-Software Control